## RFx 3000016249
Attachment D - Price Sheet
Solicitation 3000022790 - Supplemental Emergency Contract for Medical Oxygen Cylinders and Supplies
Line Item # | Item Description | Qty | Unit of Measurement | Unit Price
1 | ** PURCHASE** OXYGEN CYLINDER, Aluminum or Steel, Full, SIZE C, CGA-870 POST VALVE, 246 LITERS/8.7 CU FT 2015 PSI, DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): ____1 ____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
2 | **FILL/REFILL** OXYGEN CYLINDER, Aluminum or Steel, SIZE C, CGA-870 POST VALVE, 246 LITERS/8.7 CU FT 2015 PSI, DOT APPROVED. Type of Cylinder Bidding ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
3 | ** RENTAL** OXYGEN CYLINDER, Aluminum or Steel, Full, SIZE C, CGA-870 POST VALVE, 246 LITERS/8.7 CU FT 2015 PSI, DOT APPROVED. Type of Cylinder Bidding ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): _____1 _____2 ______3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Month
4 | COST TO REPLACE OXYGEN CYLINDER IF LOST AND/OR NOT RECOVERED. LINE ITEM BID FOR LINE ITEM #3, SIZE C. | 1 | Each
5 | **PURCHASE** OXYGEN CYLINDER, Aluminum or Steel, Full, SIZE D, CGA-870 POST VALVE, 425 LITERS/15.0 CU FT, 2015 PSI, DOT APPROVED. Type of Cylinder Bidding ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): ____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
6 | **FILL/REFILL** OXYGEN CYLINDER, Aluminum or Steel, SIZE D, CGA-870 POST VALVE, 425 LITERS/15.0 CU FT 2015 PSI, DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
7 | ** RENTAL** OXYGEN CYLINDER, Aluminum or Steel, FULL SIZE D, CGA-870 POST VALVE, 425 LITERS/15.0 CU FT, 2015 PSI, DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): ______1 _____2 _____3 ______4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Month
8 | COST TO REPLACE OXYGEN CYLINDER IF LOST AND/OR NOT RECOVERED. LINE ITEM BID FOR LINE ITEM #7, SIZE D. | 1 | Each
9 | ** PURCHASE** OXYGEN CYLINDER, Aluminum or Steel, Full, SIZE E, CGA-870 POST VALVE, 679 LITERS/24.0 CU FT . 2015 PSI, DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
10 | **FILL/REFILL** OXYGEN CYLINDER, Aluminum or Steel, SIZE E, CGA-870 POST VALVE, 679 LITERS/24.0 CU. FT. 2015 PSI DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): ____1 _____2 ____3 ____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
11 | **RENTAL** OXYGEN CYLINDER, Aluminum or Steel, Full, SIZE E, CGA-870 POST VALVE, 679 LITERS/24.0 CU. FT., 2015 PSI DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): ____1 ____2 ____3 ____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Month
12 | COST TO REPLACE OXYGEN CYLINDER IF LOST AND/OR NOT RECOVERED. LINE ITEM BID FOR LINE ITEM #11, SIZE E. | 1 | Each
13 | **PURCHASE** OXYGEN CYLINDER, Aluminum or Steel, Full, SIZE H, CGA-540 VALVE WITH CAP, 7970 LITERS, 2200 PSI, DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): _____1 ____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
14 | **FILL/REFILL** OXYGEN CYLINDER, Aluminum or Steel, SIZE H, CGA-540 VALVE WITH CAP, 7970 LITERS, 2200 PSI, DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 ____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
15 | **RENTAL** OXYGEN CYLINDER, Aluminum or Steel, Full, SIZE H, CGA-540 VALVE WITH CAP, 7970 LITERS, 2200 PSI, DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Month
16 | COST TO REPLACE OXYGEN CYLINDER IF LOST AND/OR NOT RECOVERED. LINE ITEM BID FOR LINE ITEM #15, SIZE H. | 1 | Each
17 | **PURCHASE** OXYGEN CYLINDER, Aluminum or Steel, FULL SIZE M, CGA-540 SPIGOT VALVE 3454 LITERS/122.0 CU. FT., 2216 PSI, DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): ____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
18 | **FILL/REFILL** OXYGEN CYLINDER, Aluminum or Steel, SIZE M, CGA-540 SPIGOT VALVE 3454 LITERS/122.0 CU. FT., 2216 PSI, DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): ____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
19 | **RENTAL** OXYGEN CYLINDER, Aluminum or Steel, Full, SIZE M, CGA-540 SPIGOT VALVE 3454 LITERS/122.0 CU. FT., 2216 PSI, DOT APPROVED. Type of Cylinder Bidding: ______Aluminum _____Steel REGION(S) Bidding (Check all that are applicable): _____1 _____2 ______3 ______4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Month
20 | COST TO REPLACE OXYGEN CYLINDER IF LOST AND/OR NOT RECOVERED. LINE ITEM BID FOR LINE ITEM #19, SIZE M. | 1 | Each
21 | **PURCHASE** STRAPS: ALL ASSEMBLIES WITH 10,000 LB CAPACITY AND INCLUDES TWO WIDE HEAVY DUTY LOW STRETCH WEBBING WITH 12' LENGTH. END TYPE: EACH END HAS 12' OF 5/16 CHAIN AND GRAB HOOK, DOT APPROVED. REGION(S) Bidding (Check all that are applicable): _____1 _____2 ______3 ______4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
22 | **PURCHASE** CYLINDER CART DOLLY: HOLDS (1) M60, M, H OR T SIZE CYLINDER, APPROX. DIMENSIONS AND WEIGHT: HEIGHT: 48, WIDTH: 18, DEPTH: 18 WEIGHT: 25 LBS., CASTER DIA: 3" - 5" SWIVEL, W/BRAKE, RETAINING SAFETY CHAIN. REGION(S) Bidding (Check all that are applicable): _____1 _____2 ____3 ____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
23 | **PURCHASE** OXYGEN TRANSPORT CART: CYLINDER CAPACITY: 40 CYLINDERS, CYLINDER SIZE D AND/OR E, IRON STEEL FRAME, REMOVABLE STEEL DOORS, REMOVABLE HANDLES, *BUMPER GUARD KIT*, LOCKING CASTERS, DOT APPROVED. REGION(S) Bidding (Check all that are applicable): ____1 ____2 ____3 ____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
24 | **PURCHASE** OXYGEN TRANSPORT CART: CYLINDER CAPACITY 40 CYLINDERS, LOCK TOP, CYLINDER SIZE D AND/OR E, IRON STEEL FRAME, REMOVEABLE HANDLES, LOCKING CASTERS, DOT APPROVED. REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
25 | **PURCHASE** OXYGEN TRANSPORT PALLET: METAL PALLET WITH METAL RAILINGS, APPROX. PALLET DIMENSIONS 29 1/2 X 72 X 36, DOT APPROVED. REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 ____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Number if Applicable): ____________________________________________ | 1 | Each
26 | **PURCHASE** OXYGEN CONCENTRATOR: with Oxygen Monitor; Maximum Sound Level: 55 dB; Maximum Electrical: 120VAC Input Voltage, 60Hz; Minimum Oxygen Output Concentration Levels: 87% to 96% at 0.5-5L/min; Maximum Outlet Pressure: 5 PSI (+/- 0.5 PSI); Minimum Operating Flow Rate: 0.5 to 5 Liters per minute; Must be capable of operating up to 24 hours per day; Must include the following safety features: Battery Free Power Loss Alarm, Low Pressure Alarm w/compressor shutdown, High Pressure Alarm w/compressor shutdown, High Temperature Compressor Shutdown, Current Overload or Line Surge Shutdown, Oxygen System Possible Obstruction Alert; Maximum Weight: 55 lbs; Cabinet must be Impact Resistant, Flame-Retardant Plastic that conforms to UL 94-V0; To include all of the following: Cabinet Filter, Outlet HEPA Filter, Compressor Inlet Filter, a Minimum 7 ft cannula with a maximum 50 ft of Crush-Proof Tubing; Must have a minimum of two wheels and a minimum of one handle; Invacare Perfecto2, Model # IRC 5P02 or equal REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Item Number if Applicable):______________________________________________ | 1 | Each
27 | **RENTAL** OXYGEN CONCENTRATOR with Oxygen Monitor; Maximum Sound Level: 55 dB; Maximum Electrical: 120VAC Input Voltage, 60Hz; Minimum Oxygen Output Concentration Levels: 87% to 96% at 0.5-5L/min; Maximum Outlet Pressure: 5 PSI (+/- 0.5 PSI); Minimum Operating Flow Rate: 0.5 to 5 Liters per minute; Must be capable of operating up to 24 hours per day; Must include the following safety features: Battery Free Power Loss Alarm, Low Pressure Alarm w/compressor shutdown, High Pressure Alarm w/compressor shutdown, High Temperature Compressor Shutdown, Current Overload or Line Surge Shutdown, Oxygen System Possible Obstruction Alert; Maximum Weight: 55 lbs; Cabinet must be Impact Resistant, Flame-Retardant Plastic that conforms to UL 94-V0; To include all of the following: Cabinet Filter, Outlet HEPA Filter, Compressor Inlet Filter, a Minimum 7 ft cannula with a maximum 50 ft of Crush-Proof Tubing; Must have a minimum of two wheels and a minimum of one handle; Invacare Perfecto2, Model # IRC 5P02 or equal REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Indicate whether the rental price being bid is: _____ per month ______ per week ______ per day Specify Brand (& Item Number if Applicable):___________________________________________ | 1 | Each
28 | **PURCHASE** OXYGEN REGULATOR for use with Cylinders having a CGA 540 Inlet Valve; DISS Outlet; Flow Range up to 15 LPM; Maximum Inlet Pressure: 3000 PSI; Two gauges with minimum 2" diameter: Supply Pressure Gauge up to 4000 PSI and Delivery Pressure Gauge up to 15 LPM; Sintered bronze filters in the Inlet port and Outlet Fitting; Single Stage Metering; Must have an adjustable knob for regulating flow rate; All Brass Core; Harris Products Group, Part # 3500619 or equal REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Item Number if Applicable): _____________________________________________ | 1 | Each
29 | **PURCHASE** OXYGEN REGULATOR for use with Cylinders having a CGA 870 Inlet Valve; DISS Outlet; Flow Ranges up to 15 LPM; Maximum Inlet Pressure: 3000 PSI; Two gauges with minimum 1-1/2" diameter: Supply Pressure Gauge up to 4000 PSI and Delivery Pressure Gauge up to 15 LPM; Sintered bronze filters in the Inlet Port and Outlet Fitting; Single Stage Metering; Must have an adjustable knob for regulating flow rate; All brass core; Harris Products Group, Part # 3500603 or equal REGION(S) Bidding (Check all that are applicable): ______1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Item Number if applicable): ___________________________________________ | 1 | Each
30 | **PURCHASE** SWIVEL-STYLE CHRISTMAS TREE ADAPTOR Nut and Nipple to Tapered Barb; Must be DISS threaded; Plastic Swivel Connection; Maximum Pressure: 200 PSI; 9/16"-18 UNF; Priced per Each Medica Medical Products, Part # M24-45P or equal REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 ______5 ______6 _____7 _____8 _____9 Specify Brand (&Item Number if applicable): _______________________________________________ Specify Package Size (if applicable): ______________________________________________________ | 1 | Each
31 | **PURCHASE** LARGE EMERGENCY/MASS CASUALTY OXYGEN DISTRIBUTION SYSTEM Eight Constant Flow Selector Valves, each with adjustable flow rates of 2, 4, 6, 10, and 15 LPM; one standard male oxygen DISS check valve fitting; Must have a 25 micron stainless steel mesh screen on each selector valve; Nominal Inlet Pressure: 50 PSIG; Minimum Hose Length: 10 ft.; Maximum Weight: 20lbs; Body made of anodized aluminum; Fittings made of plated brass; Includes heavy duty polyethylene hard carrying case with handle (capable of fitting all components), eight bubble humidifiers, eight barbed fittings, and eight adult size, single use masks, each with a minimum 7 ft. of oxygen supply tubing; Allied Healthcare Life Support Products, Multilator, Item # L405 or equal REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Item Number if Applicable): ______________________________________________ | 1 | Each
32 | **PURCHASE** SMALL EMERGENCY/MASS CASUALTY OXYGEN DISTRIBUTION SYSTEM 4" Anodized Aluminum Manifold, with Five Oxygen DISS check valves, capable of functioning simultaneously, attachable to an oxygen source; 50 PSIG Outlets; Supply Pressure: 40 to 90 PSIG; Flow Rates: 4, 6, 10, or 15 LPM, depending upon the metered orifice used; Fittings made of plated brass; Orifices made of stainless steel; Includes five screw-on 15 LPM orifices, and zippered nylon carry case (capable of fitting all components); Capable of being attached in tandem if more outlets are needed; Allied Healthcare Life Support Products, Minilator, Item # L419-050 or equal REGION(S) Bidding (Check all that are applicable): _____1 _____2 _____3 _____4 _____5 _____6 _____7 _____8 _____9 Specify Brand (& Item Number if Applicable):______________________________________________ | 1 | Each